
NORTHERN TERRITORY OF AUSTRALIA
radiation (safety control) act

As in force at 1 January 1984
Table of provisions
Part I
Preliminary
1
Short title
1
2
Commencement
1
3
Definitions
1
4
Application
4
5
Exemptions
5
Part II
Administration
6
Administration of Act
5
7
Delegation
5
8
Inspectors and their powers
5
9
Seizure of articles
7
Part III
Radiation safety
10
Exemptions from Part
7
11
Licences
8
12
Conditions and terms of licences
8
13
Records to be kept
9
14
Duty of licensee to ensure compliance
9
15
Obligations of a licensee in relation to irradiating apparatus, &c.
9
16
Medical examinations
10
17
Duties of radiation safety officer
10
18
Duties of persons on licensed premises
11
19
Duties of licensees
12
20
Measurement of ionizing radiation on premises
12
21
Use of measuring instruments, &c.
12
22
Disposal of radioactive substances
13
23
Maximum doses of radiation
13
24
Persons receiving dose exceeding maximum not to be employed
13
25
Precaution to be taken against receipt of excessive doses of radiation
13
26
Records to be kept
14
27
Excessive doses to be reported
14
28
Matters to be considered in computing maximum permissible dose
14
Part IV
Registration of irradiating apparatus
29
Irradiating apparatus to be registered
15
30
Conditions under which apparatus registered
15
31
Certificate to be issued and displayed
15
32
Apparatus not to be altered or modified
16
33
Inspection
16
34
Limitation of use of irradiating apparatus
16
Part V
Transport of radioactive substances
35
Exemptions
16
36
Conditions for transport of radioactive substances
17
37
Classification of packages
17
38
Materials enclosed in Type A or Type B packages
17
39
Materials to be enclosed in Type B packages
18
40
Packages to be tested
19
41
Packages not to be altered
19
42
Other articles not to be carried with radioactive substances
19
43
Conditions for transport
19
44
Sign to be displayed
20
45
Standards for conveyance of packages in single container
20
46
Liquid pyrophoric substances not to be carried
21
47
Leaking packages not to be used
21
48
Check of vehicles for radioactive contamination
21
49
Warning to be given
22
50
Procedure in case of accident
22
51
Isolation of site of accident
22
52
Fissile substances not to be carried
22
Part VI
Stores
53
Approved stores
23
Part VII
Miscellaneous
54
Offences
23
54A
Regulatory offences
23
55
Evidentiary certificates
24
56
Regulations
24
Schedule 1
Maximum permissible concentration for discharge of radioactive substances in air and water
Schedule 2
Radiation warning labels and notices
Schedule 3
Schedule 4
Classification of radioactive substances into groups for transport purposes
Schedule 5
Maximum amounts of radioactive substances exempted from this Act
Schedule 6
Specifications for packages
ENDNOTES


NORTHERN TERRITORY OF AUSTRALIA
____________________

As in force at 1 January 1984
____________________

radiation (safety control) act

An Act relating to the control, regulation, possession, use and transport of radioactive substances and irradiating apparatus

Part I
Preliminary
1
Short title
This Act may be cited as the Radiation (Safety Control) Act.

2
Commencement
This Act shall commence on a date to be fixed by the Administrator by notice in the Gazette.

3
Definitions

(1)
In this Act, unless the contrary intention appears:
absorbed dose of any radiation means the amount of energy imparted to matter by ionizing particles per unit of mass of irradiated material at a particular place and is expressed in rads.
alpha particle means a corpuscle consisting of the nucleus of a helium atom emitted by a radioactive atomic nucleus.
beta particle means a high speed corpuscle having the same mass or charge as the electron or positron and originating from nuclear disintegration.
Chief Medical Officer means the Chief Medical Officer appointed under the Public Health Act.
container means a freight container.
curie is the measure of radioactivity and one curie represents an activity of 3.7 x 1010 nuclear transformations per second.
dose equivalent is the quantity, measured in rems, which expresses on a common scale for all ionizing radiation the irradiation incurred by exposed matter.
dose-rate means the absorbed dose received per unit of time.
electron means a high speed subatomic particle or corpuscle which is negatively charged and which originates from an extra-nuclear region of the atom.
exempted substance means a substance exempted in pursuance of section 35.
gamma ray means an electromagnetic ionizing radiation which originates from the nucleus of an atom but which is not otherwise distinguishable from X-rays.
handle includes load, unload, discharge, stack, stow, store, transport and any operation incidental to or arising out of those operations.
inspector means an inspector appointed under section 8.
irradiating apparatus means an instrument or apparatus that is capable of emitting radiation not originating from a radioactive substance in or on the instrument or apparatus, but does not include:
(a)
an instrument or apparatus from which the dose rate to an organ or tissue of a person when situated at a distance 0.1 of a metre from the external surface of the instrument or apparatus does not exceed 0.1 millirem per hour;

(b)
an instrument or apparatus in which electrons are accelerated to an energy not exceeding 5000 electron volts; or

(c)
a television receiver from which the dose rate to an organ or tissue of a person when situated at a distance of 5 centimetres from the external surface of the receiver does not exceed 0.5 of a millirem per hour.
licensed premises means the premises described in a licence issued under this Act.
licensee means a person licensed under this Act.
maximum permissible dose, in relation to a person or a part of the body of a person exposed to radiation, means:
(a)
for any quarter of a calendar year – the dose equivalent of the radiation specified in the third column of Part I of Schedule 3 received through the organs or parts of his body specified in the second column of that Part; and

(b)
for the period of a calendar year – the dose equivalent of the radiation specified in the third column of Part II of Schedule 3 received through the organs or parts of his body specified in the second column of that Part.
maximum permissible concentration means a concentration of a radioactive substance specified in the first column of Schedule 1:
(a)
in air, being that specified in the second column of that Schedule; or

(b)
in water, being that specified in the third column of that Schedule,

for the radioactive substance.
neutron means a radioactive corpuscle which has no electric charge and has a mass slightly greater than the proton.
nuclide means a variety of an atom characterized by a specific atomic number and a specific mass number.
package includes a pack, packet, parcel, carton, box or receptacle of any kind that contains a radioactive substance.
place and premises includes buildings, ships, aircraft, vehicles and any other premises on land or water and all other land, vacant or otherwise whether public or private.
positron means a high speed, subatomic particle or corpuscle which is positively charged and which originates from an extra nuclear region of the atom.
proton means a nuclear particle of unit mass number having a charge equal to and opposite to that of an electron.
rad is the unit of measurement of the absorbed dose and one rad represents the absorbed dose received when energy of 0.01 joule is imparted to one kilogram of irradiated matter.
radiation means:

(a)
electromagnetic radiation, being X-rays, gamma rays, photons or quanta; or

(b)
corpuscular radiation being alpha particles, beta particles, electrons, protons, neutrons and heavy particles capable of causing ionization of matter on which it impinges.
radiation hazard means a thing or situation that creates a danger to the health of a person and that arises from exposure of ionizing radiation because of radiation from an external source or from within the body.
Radiation Safety Officer means a person appointed to be a Radiation Safety Officer under section 15.
radiation worker means a person who receives or is likely to receive radiation in the course of his employment whether from a radioactive substance or from irradiating apparatus.
radioactive contamination means the lodgement, attachment or incorporation of a radioactive substance, on, to or in an organ or tissue of a person or on or to any other material or substance.
radioactive substance means a substance which consists of or contains radioactive nuclides whether natural or artificial.
radioactivity means the spontaneous disintegration of an unstable nuclide with the emission of a particle or proton to form a different nuclide.
rem is the unit measurement of dose equivalent and one rem is the dose equivalent when irradiated matter receives an absorbed dose of one rad.
X-ray means electromagnetic, ionizing radiation which originates from the field outside the nucleus of the atom and resulting from the loss of energy of charged particles.


(2)
Where in this Act a unit of measurement is used coupled with the prefix "milli" or "micro", the prefix signifies that the measurement referred to is one-thousandth or one-millionth respectively of the unit of measurement.

4
Application

(1)
Subject to subsection (2), this Act is binding upon all persons in the Territory whether or not they are in the service of any government or any governmental authority.


(2)
For 21 days immediately following the date of commencement of this Act, this Act shall not apply to a person who, on that date of commencement has in his possession or is using irradiating apparatus or a radioactive substance.

5
Exemptions

(1)
A radioactive substance specified in Schedule 5 is exempt from the operation of this Act if its radioactivity does not exceed the maximum specified for it in that Schedule.


(2)
Nothing in this Act applies to the mining, production, possession, treatment, handling, sale, use or disposal of uranium ores or uranium oxide (U3O8).


(3)
The Chief Medical Officer, by notice in the Gazette, may exempt from the operation of this Act any radioactive substance or irradiating apparatus which he is satisfied does not produce a significant radiation hazard.

Part II
Administration
6
Administration of Act
The Chief Medical Officer is, in the performance of his duties and functions and the exercise of his powers under this Act, subject to the direction and control of the Minister.

7
Delegation

(1)
The Chief Medical Officer may, by instrument in writing, delegate to a person all or any of his powers, functions and authorities under this Act (except this power of delegation) in relation to a matter or class of matters or to a district or part of the Territory so that the delegated powers, functions and authorities may be exercised by the delegate with respect to the matter or class of matters or the district or part of the Territory specified in the instrument of delegation.


(2)
A delegation under subsection (1) is revocable in writing at will and does not prevent the exercise of a power, function or authority by the Chief Medical Officer.

8
Inspectors and their powers

(1)
The Chief Medical Officer may appoint such inspectors as he thinks necessary for the purposes of the administration of this Act.


(2)
An inspector may:
(a)
with the consent, expressed or implied, of the owner or occupier of premises enter those premises and search for and examine:
(i)
registers and licences kept or held under this Act;

(ii)
books, records and documents relating to irradiating apparatus or equipment used in connection with that apparatus and any instruments, devices or accessories for measuring or detecting radiation or radioactive substances; and

(iii)
packages containing or which he believes, on reasonable grounds, contain irradiating apparatus or a radioactive substance;

(b)
use or install any equipment or measuring device for sampling, measuring or analyzing any substance; or

(c)
make copies of or reproduce writings, books, records or other recorded information whether the information be recorded in writing or in some other form.


(3)
An inspector, without the consent of any person but in pursuance of a search warrant, may enter any premises or place and do all such acts as are permitted to be done by subsection (2) with the consent of the owner or occupier of the premises.


(4)
Where there are reasonable grounds for an inspector to believe that a radiation hazard exists and that that hazard is such that action under this section is urgent, he may do all such acts as are permitted to be done under this section in pursuance of a search warrant without the consent of the owner or occupier and without obtaining a search warrant and he shall report any action taken under this subsection to the Chief Medical Officer within 48 hours.


(5)
If a Justice of the Peace is satisfied by information on oath that there are reasonable grounds for suspecting that there is in any place any thing:
(a)
with respect to which an offence against this Act has been or is suspected on reasonable grounds to have been committed;

(b)
as to which there are reasonable grounds for suspecting that it will afford evidence of that offence; or

(c)
as to which there are reasonable grounds for believing that it is intended to be used for the purpose of committing such an offence against this Act, he may grant a search warrant authorizing an inspector or member of the Police Force named in the warrant, with such assistance as he thinks necessary, to enter at any time a place described in the warrant and to seize any such thing that he may find in that place.

9
Seizure of articles

(1)
When an inspector has reasonable grounds to believe that a radiation hazard exists or that an offence has been or is being committed against this Act he may:
(a)
seize and remove any thing constituting or contributing to a health hazard;

(b)
seize any thing likely to afford evidence of an offence against this Act; and

(c)
if a thing seized is, in the opinion of the inspector, too dangerous or too bulky to be removed, forthwith prohibit the use, removal or alteration of, and interference or contact with, the thing by all persons for a period of 14 days by a notice in writing served on the owner or occupier of the premises and exhibited on the premises as near as possible to the thing seized.


(2)
A person shall not use, move or alter, or interfere or make contact with, a thing to which a notice under subsection (1)(c) applies.

Part III
Radiation safety
10
Exemptions from Part
This Part does not apply:
(a)
in the case of a person lawfully possessing, using or operating irradiating apparatus under and in accordance with the Radiographers Act – to or in relation to that apparatus;

(b)
in the case of a patient undergoing treatment which involves the possession of a radioactive substance – to or in relation to that treatment; or

(c)
in the case of the owner of an animal undergoing treatment which involves the retention in or on the animal of a radioactive substance – to or in relation to that treatment.

11
Licences

(1)
Subject to this section, a person shall not have in his possession, manufacture, purchase, use, sell, handle or dispose of a radioactive substance or irradiating apparatus or carry out maintenance on irradiating apparatus.


(2)
On the application of a person in the prescribed form, the Chief Medical Officer may, subject to section 12(1), grant that person a licence under this section.


(3)
A licence under this section authorizes the person to whom it is granted to do, subject to section 12(2), such of the things prohibited by subsection (1) as are specified in the licence.

12
Conditions and terms of licences

(1)
The Chief Medical Officer shall not grant a licence under section 11 unless:
(a)
the applicant is a natural person over the age of 18 years; and

(b)
the Chief Medical Officer approves the means to be used in storing the radioactive substance or irradiating apparatus.


(2)
A licence granted under section 11 may be issued subject to such conditions as the Chief Medical Officer considers necessary and as are specified in the licence.


(3)
A licence granted under section 11:
(a)
remains in force for such period not exceeding 12 months as is specified in the licence;

(b)
is renewable, on application by the licensee, for such period not exceeding 12 months as the Chief Medical Officer may specify in a written notice given to the licensee;

(c)
may be cancelled by the Chief Medical Officer:
(i)
where it has been issued subject to conditions and the licensee is in breach of any of those conditions;

(ii)
where the licensee has failed, refused or neglected to comply with this Act; or

(iii)
where the Chief Medical Officer is of the opinion that, in the interest of the safety of members of the public, the licence should be cancelled; and

(d)
shall describe the premises to which the licence applies.

13
Records to be kept
A licensee shall keep in a register a record of all irradiating apparatus and radioactive substances that come into his possession and the use to which that irradiating apparatus or those radioactive substances are put and any change in that use.

14
Duty of licensee to ensure compliance
A licensee shall take all reasonable steps to ensure that every person under his supervision complies with this Act.

15
Obligations of a licensee in relation to irradiating apparatus, &c.

(1)
A licensee who has in his possession any irradiating apparatus or radioactive substance shall:
(a)
appoint a person to be a Radiation Safety Officer in respect of the licensed premises and all irradiating apparatus and radioactive substances in his possession;

(b)
give notice to the Chief Medical Officer in writing within 24 hours:
(i)
of the first occasion upon which an irradiating apparatus or a radioactive substance comes into his possession; and

(ii)
of the appointment of a Radiation Safety Officer furnishing his own private address and telephone number and the name, private address and telephone number, if any, of the Radiation Safety Officer appointed;

(c)
forward a copy of all reports and recommendations of the Radiation Safety Officer to the Chief Medical Officer within 24 hours of the receipt of the report or the recommendation;

(d)
at least 24 hours before he employs a person as a radiation worker, notify the Chief Medical Officer of the name, qualifications (if any) and experience of the person proposed to be so employed; and

(e)
within 24 hours after a person ceases to be so employed notify the Chief Medical Officer of the cessation.


(2)
Before a Radiation Safety Officer is appointed or during the absence from licensed premises of the Radiation Safety Officer, the licensee has, with respect to the licensed premises and all irradiating apparatus and radioactive substances on those premises, all the duties, powers, liabilities and responsibilities prescribed by this Act in respect of a Radiation Safety Officer.

16
Medical examinations

(1)
The Chief Medical Officer may require a licensee, a person appointed to be a Radiation Safety Officer or a radiation worker to submit to such medical examination as the Chief Medical Officer may specify.


(2)
A person required to submit himself to a medical examination under subsection (1) shall not fail, refuse or neglect to do so.

17
Duties of radiation safety officer
A Radiation Safety Officer shall:
(a)
investigate and record all sources of ionizing radiation on premises under his control;

(b)
within 14 days of his appointment and thereafter at 6 monthly intervals prepare a report:
(i)
showing the names of all persons permitted by the licensee to be in a place where ionizing radiation exists or may, from time to time, exist;

(ii)
specifying the times of the day during which each person is likely to be in such a place and exposed to ionizing radiation;

(iii)
recommending the safe working procedures that should be adopted for work on the licensed premises in connection with irradiating apparatus or radioactive substances; and

(iv)
recommending, if necessary, the installation or use of facilities for the purpose of minimizing the absorbed dose that each person may receive;

(c)
provide a copy of each report made under paragraph (b) to each person working in the place to which his report relates who may be subjected to ionizing radiation and to the employer of each of those persons;

(d)
take all reasonable steps to ensure that all persons likely to be subjected to ionizing radiation are adequately instructed in the use of all safeguards and procedures and are supplied with such apparatus, clothing, instruments, shields, devices or accessories as are necessary for the protection of those persons from ionizing radiation;

(e)
take all reasonable steps to ensure that persons not engaged in work involving the use or handling of irradiating apparatus or radioactive substances cannot be subject to ionizing radiation (other than that naturally occurring) exceeding 10 millirems a week;

(f)
take all reasonable steps to ensure that no radioactive substance is removed from the licensed premises in contravention of this Act;

(g)
if he becomes aware of the existence of any ionizing radiation from a source not under his control – report the matter in writing immediately to the licensee and to the Chief Medical Officer;

(h)
take all reasonable steps to ensure that all persons employed on the licensed premises carry out all such procedures and do all such acts as will ensure the safe performance of their work;

(i)
keep a register of every sealed radiation source on the licensed premises;

(j)
at the cessation of work on each day, ensure that each source of ionizing radiation used on that day is accounted for and that other sources of ionizing radiation are held in a safe and secure place; and

(k)
take all reasonable steps to ensure that all apparatus, instruments, devices and accessories used for the protection of persons from ionizing radiation or for the detection and measurement of ionizing radiation, absorbed doses and dose equivalents and of radioactive contamination are maintained in good working condition and are properly used.

18
Duties of persons on licensed premises
A radiation worker shall use in the proper manner all apparatus, instruments, devices, clothing, shields and accessories supplied to him for his protection and shall observe the working procedures laid down by the Radiation Safety Officer appointed for the premises in which the worker works.

19
Duties of licensees
A licensee shall:
(a)
by means of doors, bars, locks or warning or cautionary notices, signs or lights, prohibit the access of unauthorized persons to all parts of the licensed premises in which they may be subjected to ionizing radiation;

(b)
immediately upon becoming aware that a radioactive substance in his possession or under his control has been damaged, lost or involved in an accident or fire, notify the Chief Medical Officer of the fact by telegram, telephone or personal communication and shall confirm that notification in writing as soon as is reasonably possible;

(c)
carry out all instructions that the Chief Medical Officer gives to him consequent upon a notification under paragraph (b); and

(d)
take all reasonable steps to ensure that the concentration of radioactive substances in air and water does not exceed the concentrations prescribed in Schedule 1.

20
Measurement of ionizing radiation on premises
A licensee who employs persons as radiation workers shall:
(a)
carry out, when and in the manner required by the Chief Medical Officer, measurements of ionizing radiation in and around the licensed premises and in air and water discharged from the licensed premises;

(b)
provide and maintain for each radiation worker such instruments, apparatus, devices or accessories as the Chief Medical Officer requires for the purpose of measuring the amount of ionizing radiation to which a radiation worker is or has been exposed; and

(c)
instruct those workers in the methods in which those instruments, apparatus, devices or accessories are to be used.

21
Use of measuring instruments, &c.
A radiation worker using or handling an irradiating apparatus or a radioactive substance shall, while doing so or while in the vicinity of an irradiating apparatus or a radioactive substance, carry attached to his person or his clothing such instrument, apparatus, device or accessory as shall have been provided in accordance with section 20(b).

22
Disposal of radioactive substances
A person shall not dispose of or abandon a radioactive substance without the approval of the Chief Medical Officer.

Penalty:
$5000 or imprisonment for 2 years.

23
Maximum doses of radiation
An employer or the person in charge of a place where irradiating apparatus or a radioactive substance is used shall take all reasonable steps to ensure that a person on premises under his control:
(a)
does not receive a radiation dose in excess of the maximum permissible dose; and

(b)
is not exposed to a concentration of radioactive substances in air or water in excess of the maximum permissible concentration.

24
Persons receiving dose exceeding maximum not to be employed

(1)
Where the Chief Medical Officer is satisfied that a radiation worker has been exposed to radiation in excess of the maximum permissible dose, he may direct the person by whom the worker is employed to cease to employ him, for such period as the Chief Medical Officer specifies, in any employment in which he will or may be exposed to ionizing radiation.


(2)
An employer shall not fail, refuse or neglect to comply with a direction given under subsection (1).

25
Precaution to be taken against receipt of excessive doses of radiation

(1)
Where there is cause to believe that there will be a serious risk to the health of a radiation worker if he continues to be exposed to ionizing radiation, the Chief Medical Officer may direct his employer to cease to employ him in any work which may expose him to further ionizing radiation.


(2)
An employer shall not fail, refuse or neglect to comply with a direction given under subsection (1).

26
Records to be kept
An employer of a radiation worker or the person in charge of a place where radiation workers are employed shall keep a record in a form approved by the Chief Medical Officer showing:
(a)
the full name, address, age and sex of each radiation worker;

(b)
the date of commencement of employment;

(c)
the dates from which and the periods during which the person has or may have been exposed to ionizing radiation;

(d)
the details of all dose assessments for that person; and

(e)
all known facts relating to any accidental dose of ionizing radiation that may have been received by that person.

27
Excessive doses to be reported
An employer of a radiation worker and a person in charge of a place in which irradiating apparatus or a radioactive substance is used:
(a)
who has cause to suspect that a person has received (other than as a patient undergoing treatment) a dose equivalent in excess of 200 millirems per week or that some unusual occurrence has taken place in or about a source of ionizing radiation; or

(b)
who becomes aware that a personal monitoring device has recorded in respect of a person a dose equivalent exceeding 200 millirems per week,

shall report the fact to the Chief Medical Officer forthwith.

28
Matters to be considered in computing maximum permissible dose

(1)
In calculating the maximum permissible dose of ionizing radiation for the purposes of this Act, doses:
(a)
received by an organ or tissue of a person from any source of ionizing radiation as a consequence of the employment of that person shall be counted;

(b)
received by a person in undergoing radiological procedures or ionizing radiation received by him as the result of exposure to naturally occurring radiation shall not be counted; and

(c)
received in any quarter of a calendar year shall be counted in the dosage for the calendar year of which that quarter forms part.


(2)
If the dose to the whole body, gonads or red bone marrow accumulated by a person in the course of employment is not known for any period, it shall be assumed that he has received, in each calendar year of the period, the relevant maximum permissible dose for those parts specified in Part II of Schedule 3.

Part IV
Registration of irradiating apparatus
29
Irradiating apparatus to be registered
A person shall not use irradiating apparatus unless the apparatus is registered under this Part.

30
Conditions under which apparatus registered

(1)
A person who is in possession of irradiating apparatus shall forthwith make application to the Chief Medical Officer for the registration of that irradiating apparatus.


(2)
If the Chief Medical Officer is satisfied that:
(a)
the equipment is in a safe and good operating condition;

(b)
the equipment is suitable for the use proposed;

(c)
the location and installation of the equipment are appropriate; and

(d)
the equipment is adequately protected and, if shielding of the equipment is necessary, that it is properly shielded,

he shall register the irradiating apparatus.

31
Certificate to be issued and displayed

(1)
the Chief Medical Officer shall issue a certificate of registration in respect of each item of irradiating apparatus that is registered.


(2)
A person to whom a certificate of registration is issued shall display the certificate in a prominent position on or near the irradiating apparatus to which it relates.

32
Apparatus not to be altered or modified
A person shall not alter or modify any registered irradiating apparatus without the approval of the Chief Medical Officer.

33
Inspection

(1)
An inspector shall regularly inspect each item of registered irradiating apparatus and report to the Chief Medical Officer on that inspection.


(2)
If he is satisfied upon the report of an inspector that any registered irradiating apparatus is in a dangerous condition or requires repair or modification, the Chief Medical Officer may cancel the registration of the apparatus and require the person in possession of the apparatus to deliver to the Chief Medical Officer the certificate of registration forthwith.


(3)
A person shall not fail, refuse or neglect to comply with a requirement of the Chief Medical Officer under subsection (2).

34
Limitation of use of irradiating apparatus

(1)
A certificate of registration may be issued subject to such conditions as the Chief Medical Officer considers necessary and as are endorsed on the certificate.


(2)
Where a certificate of registration is endorsed under subsection (1), a person shall:
(a)
not use the irradiating apparatus for a purpose other than the purpose specified in the certificate; and

(b)
observe the conditions endorsed on the certificate.

Part V
Transport of radioactive substances
35
Exemptions

(1)
The Chief Medical Officer, by notice published in the Gazette, may exempt from the provisions of this Part a radioactive substance of a weight or activity specified in the notice.


(2)
A radioactive substance referred to in a notice published in pursuance of subsection (1), is exempt from this Part.

36
Conditions for transport of radioactive substances
A person shall not transport a radioactive substance unless:
(a)
it is an exempted substance;

(b)
it is contained in a package that complies with this Part and with Schedule 6 and is transported in a manner prescribed in this Part; or

(c)
it is transported in accordance with the requirements of Safety Series No. 6 Regulations for the Safe Transport of Radioactive Materials 1973 Revised Edition published by the International Atomic Energy Agency.

Penalty:
$1,000 or imprisonment for 12 months.

37
Classification of packages

(1)
For the purposes of this Part, packages are described as Type A packages and Type B packages.


(2)
Type A packages shall comply with the specifications prescribed in Schedule 6.


(3)
Type B packages:
(a)
shall comply with the specifications prescribed in Schedule 6 and with such additional requirements as may be prescribed or specified by the Chief Medical Officer; and

(b)
shall have their outer receptacles resistant  to fire and water.

38
Materials enclosed in Type A or Type B packages
A person shall not transport:

(a)
a radioactive substance specified in the first column of Schedule 4 and classified in the second column of that Schedule opposite that substance as being:
(i)
in Group I if it has a maximum radioactivity of 100 microcuries;

(ii)
in Group II if it has a maximum radioactivity of 10 millicuries; or

(iii)
in Group III if it has a maximum radioactivity of 2 curies; or

(b)
a radioactive substance in the form of a non-friable massive solid that:
(i)
has a melting point exceeding 537 degrees Celsius;

(ii)
is not soluble in water;

(iii)
does not react with air or water; and

(iv)
has a maximum radioactivity of 20 curies,

except in a Type A package or a Type B package.

Penalty:
$2,000 or imprisonment for 2 years.

39
Materials to be enclosed in Type B packages
A person shall not transport:
(a)
a radioactive substance specified in the first column of Schedule 4 and classified in the second column of that Schedule opposite the substance specified in the first column as being:
(i)
in Group I if it has a maximum radioactivity of 20 curies;

(ii)
in Group II if it has a maximum radioactivity of 20 curies; or

(iii)
in Group III if it has a maximum radioactivity of 200 curies; or

(b)
a radioactive substance in the form of a non-friable massive solid that:
(i)
has a melting point exceeding 537 degrees Celsius;

(ii)
is not soluble in water;

(iii)
does not react with air or water; and

(iv)
has a maximum radioactivity of 2000 curies,

except in a Type B package.

40
Packages to be tested
A person shall not carry a package unless it has been subjected to such tests as are prescribed or, failing prescription, as are set out in a notice in writing served by the Chief Medical Officer upon the person licensed to transport the substance in the package.

41
Packages not to be altered
A person shall not alter or modify a package during transport.

42
Other articles not to be carried with radioactive substances
A person shall not include in a package any other article or thing except such articles, things or documents as are necessary for the use or identification of the radioactive substances.

43
Conditions for transport

(1)
A person in charge of a vehicle, vessel or aircraft shall ensure that a radioactive substance in the course of transport is kept in a place which is separate and shielded from:
(a)
living or sleeping accommodation;

(b)
regularly occupied working places;

(c)
all places which may be at any time occupied by passengers or members of the public; and

(d)
all places in which there is stowed an explosive or corrosive material or a material which is capable of spontaneous combustion.


(2)
A person shall not load a package into the same vehicle or aircraft or hold, compartment or deck area of a vessel as an explosive or corrosive material or a material which is capable of spontaneous combustion which might damage the package in the event of an accident.


(3)
A person shall not transport or permit the transportation of a package which is required under Schedule 6 to be labelled Category II or Category III in a compartment which is occupied by passengers.

44
Sign to be displayed
The person in charge of a vehicle transporting a package shall cause a sign:
(a)
conforming with figure 5 in Schedule 2; and

(b)
complying with the colour and other requirements indicated in that Schedule for that figure,

to be displayed on the 2 sides and on the rear of the vehicle so that one of those signs is clearly visible from all directions except from in front of the vehicle.

45
Standards for conveyance of packages in single container
Notwithstanding anything contained in this Act, where a load of packages is transported in one container, from one consignor having the sole use of the vehicle and in respect of which all loading and unloading is carried out by the consignor or consignee, the following levels of radiation may be observed:

(a)
1000 millirems per hour at any point on the external surface of the container if:
(i)
the vehicle is equipped with an enclosure to hold the container so that, during normal transport, access by unauthorized persons to the interior of the enclosure is prevented;

(ii)
the packages and the container are secured within the enclosure so that their position remains fixed during transport; and

(iii)
there are no loading or unloading operations between the beginning and the end of the transport;

(b)
200 millirems per hour at any point on the outer surface of the vehicle, including the upper and lower surfaces, or, in the case of an open vehicle, at any point on the verticle planes projected from the outer edges of the vehicle, on the upper surfaces of the load and on the lower external surfaces of the vehicle;

(c)
10 millirems per hour at any point 2 metres from the outer surfaces of the vehicle or, if the load is transported in an open vehicle, at any point 2 metres from the verticle planes projected from the outer edges of the vehicle; and

(d)
2 millirems per hour in any position in the vehicle occupied by persons, except with the approval of the Chief Medical Officer.

46
Liquid pyrophoric substances not to be carried
A person shall not load into an aircraft or carry by air a liquid pyrophoric radioactive substance.

47
Leaking packages not to be used

(1)
A person shall not use an empty package as a container for a radioactive substance where there are reasonable grounds for suspecting that a leakage of a radioactive substance from the package has occurred unless the package has been repaired and decontaminated.


(2)
A person having in his possession or under his control a package from which there are reasonable grounds for suspecting a leakage of a radioactive substance has occurred shall prevent other persons from having access to the package and shall notify the Chief Medical Officer forthwith of the suspected leakage.

48
Check of vehicles for radioactive contamination

(1)
A person having in his possession or under his control a vehicle or equipment used regularly for the transportation of radioactive substances shall, at least once a month, check that vehicle or equipment to determine the level of radioactive contamination.


(2)
A person having in his possession or under his control a vehicle or equipment which has been contaminated by a radioactive substance shall ensure that it is not used to transport other goods until it has been decontaminated and has been declared to be safe by a person appointed by the Chief Medical Officer for the purpose.


(3)
A person shall not for any purpose other than the transportation of a radioactive substance, use or carry a package which has contained radioactive substances unless:
(a)
all labels referring to a radioactive substance have been removed or covered; and

(b)
a person authorized by the Chief Medical Officer is satisfied that all accessible surfaces of the package have been monitored and that radiation has been reduced to a safe level.

49
Warning to be given
The driver of a vehicle transporting a radioactive substance shall ensure that there is affixed to the inside of the driver's cabin a conspicuous metal plate engraved, stamped or embossed as follows:

"WARNING: This vehicle is carrying a radioactive substance.  In case of damage to the vehicle or its contents notify the Chief Medical Officer and the consignor of any goods on the vehicle.".

50
Procedure in case of accident
Where a vehicle transporting a radioactive substance:
(a)
is involved in an accident;

(b)
is subjected to an unusual delay; or

(c)
may have been or may become contaminated by a leakage of a radioactive substance,

the driver shall:
(d)
forthwith notify the Chief Medical Officer of the occurrence;

(e)
obey such instructions as the Chief Medical Officer may give; and

(f)
take all reasonable steps to prevent access of any person not authorized by the Director to the vehicle or to a place in the vicinity of the radioactive substance.

51
Isolation of site of accident
Where an accident to a vehicle results in damage to a package, the driver of the vehicle and all members of the Police Force shall act as though radioactive contamination exists at the site of the accident and shall take steps to isolate the site and, until otherwise directed by the Chief Medical Officer, shall ensure that no person enters or remains on the site, unless with the Chief Medical Officer's authority.

52
Fissile substances not to be carried

(1)
A package containing a radioactive substance that is also a fissile substance shall not be transported, unless the transportation is authorized by the Chief Medical Officer in writing.


(2)
In this section fissile substance means plutonium 238, plutonium 239, plutonium 241, uranium 233 and uranium 235 but does not include uranium ores, non-irradiated uranium or depleted uranium or any substance obtained as a result of benefication of uranium ore in the form of oxides or uranium 238, uranium 234 or uranium 235 in any proportion in which they occur naturally.

Part VI
Stores
53
Approved stores

(1)
A person shall not use a place (other than licensed premises) to store a radioactive substance unless the Chief Medical Officer has approved of the place as a store.


(2)
The Chief Medical Officer shall not approve of a place as a store for radioactive substances unless he is satisfied that the facilities provided and the precautions taken are adequate to prevent any radiation hazard arising and to prevent radioactive substances being taken or used by unauthorized persons.


(3)
The Chief Medical Officer shall notify the officer in charge of the nearest police station and the Chief Fire Officer of each place in which radioactive sources are stored or held.


(4)
A store in which radioactive substances are held shall have displayed close to it or on it a radioactive warning sign:
(a)
conforming with figure 6 in Schedule 2; and

(b)
complying with the colour requirements specified in that Schedule for that figure.

Part VII
Miscellaneous
54
Offences
A person shall not fail, refuse or neglect to comply with this Act.

Penalty:
$500 or imprisonment for 6 months, or the greater penalty fixed by the relevant provision of the Act.

54A
Regulatory offences
An offence of failing, refusing or neglecting to comply with section 11(1), 13, 24, 25, 26, 29, 30(1), 31(2), 33(3), 44, 48, 49, 50(d) or (e), 52(1) or 53(1) or (4) is a regulatory offence.

55
Evidentiary certificates
Where, in a prosecution for an offence against this Act, it is necessary to prove:
(a)
the quantity of an absorbed dose;

(b)
the nature of a radioactive substance;

(c)
the activity of a radioactive substance; or

(d)
the strength of a dose equivalent,

a certificate purporting to be signed by the Chief Medical Officer stating any such measurement or other fact is evidence, in the absence of proof to the contrary, of the matters so certified and of the facts on which they are based.

56
Regulations
The Administrator may make regulations, not inconsistent with this Act, prescribing all matters which by this Act are required or permitted to be prescribed or which are necessary or convenient for giving effect to this Act including prescribing fees for the issue and renewal of licences under Part III and the registration of irradiating apparatus under Part IV.

Schedule 1
Maximum permissible concentration for discharge of radioactive substances in air and water
sections 3, 19 and 23

	First Column
	Second Column
	Third Column

	Radioactive substance
	Microcuries per millilitre of air
	Microcuries per millilitre of water

	A41
	5x10-7
	5x10-4

	Ag105
	1x10-5
	2

	Ag111
	3x10-5
	4

	Am241
	3x10-11
	1x10-4

	As76
	2x10-6
	0.2

	At211
	3x10-10
	2x10-6

	Au198
	1x10-7
	3x10-3

	Au199
	2.5x10-7
	7x10-3

	Ba140+La140
	6x10-8
	2x10-3

	Be7
	4x10-6
	1

	C14
	5x10-7
	3x10-3

	Ca45
	3x10-8
	5x10-4

	Cd109+Ag109m
	7x10-8
	7x10-2

	Cel144+Pr144
	7x10-9
	4x10-2

	C1 36
	4x10-7
	9x10-4

	Cm242
	2x10-10
	9x10-4

	Co60
	1x10-6
	2x10-2

	Cr51
	8x10-6
	0.5

	Cs137+Ba137m
	2x10-7
	1.5x10-3

	Cu64
	6x10-6
	8x10-2

	Eu154
	6x10-9
	3x10-2

	F18
	1x10-4
	0.9

	Fe55
	6x10-7
	4x10-3

	Fe59
	1.5x10-8
	1x10-4

	Ga72
	3x10-6
	9

	Ge71
	4x10-5
	9

	H3(HTO or T20)
	2x10-5
	0.2

	Ho166
	3x10-6
	23

	I131
	5x10-9
	3x10-5

	Ir190
	7x10-7
	1x10-2

	Ir192
	5x10-8
	9x10-4

	K42
	2x10-6
	1x10-2

	La140
	1x10-6
	1

	Lu177
	5x10-6
	24

	Mn56
	3x10-6
	0.15

	Mo99
	2x10-3
	14

	Na24
	2x10-6
	8x10-3

	Nb95
	4x10-7
	4x10-3

	Ni59
	2x10-5
	0.25

	P32
	1x10-7
	2x10-4

	Pb203
	6.5x10-6
	0.1

	Pd103+Rh103
	7x10-7
	1x10-2

	Pm147
	2x10-7
	1

	Po210 (soluble)
	2x10-10
	3x10-5

	Po210 (insoluble)
	7x10-11
	-

	Pr143
	7.5x10-7
	0.4

	Pu239 (soluble)
	2x10-12
	1.5x10-6

	Pu239 (insoluble)
	2x10-12
	-

	Ra226+½dr
	8x10-12
	4x10-8

	Rb86
	4x10-7
	3x10-3

	Re183
	8x10-6
	8x10-2

	Rh105
	1x10-6
	1.5x10-2

	Rn222+dr
	1x10-7
	2x10-6

	Ru106+Rh106
	3x10-8
	0.1

	S35
	1x10-6
	5x10-3

	Sc46
	7x10-8
	0.4

	Sm151
	1x10-8
	0.2

	Sn113
	6x10-7
	0.2

	Sr89
	2x10-8
	7x10-5

	Sr90+Y90
	2x10-10
	8x10-7

	Tc96
	3x10-6
	3x10-2

	Te127
	1x10-7
	3x10-2

	Te129
	4x10-8
	1x10-2

	Th234
	6x10-7
	3

	Th-natural (insoluble)
	3x10-11
	-

	Th-natural
	3x10-11
	4x10-7

	Tm170
	5x10-8
	0.25

	U233 (soluble)
	1x10-10
	1.5x10-4

	U233 (insoluble)
	1.6x10-11
	-

	U-natural (soluble)
	1.7x10-11
	7x10-5

	U-natural (insoluble)
	1.7x10-11
	-

	V48
	1x10-6
	0.5

	Xe133
	4x10-6
	4x10-3

	Xe135
	2x10-6
	1x10-3

	Y91
	4x10-8
	0.2

	Zn65
	2x10-6
	6x10-2

	All other emitters of beta or gamma radiation
	
1x10-9
	
1x10-7

	All other emitters of alpha radiation
	
5x10-12
	
1x10-7


NOTES

(1)
The figures given in the Second and Third Column of this Schedule apply to continuous exposures for 24 hours a day.  Where the exposure is incurred only during a work day of 8 hours the values in the Second Column may be multiplied by 3.

(2)
In this Schedule the expression dr means daughter products of the radioactive substance preceding the expression.

Schedule 2
Radiation warning labels and notices
sections 44 and 53

1.
Radiation warning labels shall be printed in black letters on a white or yellow background (according to the appropriate category).

2.
A source container shall be labelled as follows:

"CAUTION

RADIATION GAUGE ENCLOSING RADIOACTIVE SOURCE"


and the appropriate radiation symbol shall be displayed.

3.
A store for radioactive substance shall be labelled as follows:

"CAUTION

STORE FOR RADIOACTIVE SUBSTANCE"


and the appropriate radiation symbol shall be displayed.

Figure 1.

Basic trefoil symbol with proportions based on a central circle of radius 4 millimetres.

(See printed copy of Act)

Figure 2.

Category I – White label with lettering and trefoil black and category sign "I" denoted in red.

(See printed copy of Act)

Figure 3.

Category II – Yellow label with lettering and trefoil black and category sign "II" denoted in red.

(See printed copy of Act)

Figure 4.

Category III – Yellow label with lettering and trefoil black and category sign "III" denoted in red.

(See printed copy of Act)

Figure 5.

Placard for vehicles

The overall shape of the placard may be diamond or square as indicated by the broken lines.  Minimum dimensions are given; when larger dimensions are used the relative proportions must be maintained.

A contrasting colour other than white may be used for the background.

The lower half of the placard may be used for other hazard identification codes, letters, &c., as may be required by any law in the Territory or by conventions relating to emergency incident information systems for all classifications of dangerous goods.

(See printed copy of Act)

Figure 6.

Colours:  Background yellow lettering and trefoil black.

(See printed copy of Act)

Schedule 3
sections 3, 23, 24 and 28

PART I

MAXIMUM PERMISSIBLE DOSE ACCUMULATED IN ANY QUARTER OF A CALENDAR YEAR

	First Column
	Second Column
	Third Column

	Item No.
	Organ or part of the body
	Dose

	1.
	Gonads
	3 rems

	2.
	Red bone marrow
	3 rems

	3.
	Whole body radiation exposure
	3 rems

	
	
	

	4.
	Skin
	15 rems

	5.
	Bone
	15 rems

	6.
	Thyroid
	15 rems

	
	
	

	7.
	Hands and forearms
	40 rems

	8.
	Feet and ankles
	40 rems

	
	
	

	9.
	Any other single organ (including lens of the eye)
	
8 rems

	10.
	Abdomen of women of reproductive capacity
	
1.3 rems


PART II

MAXIMUM PERMISSIBLE DOSE ACCUMULATED IN ONE CALENDAR YEAR

	First Column
	Second Column
	Third Column

	Item No.
	Organ or part of the body
	Dose

	1.
	Gonads
	5 rems

	2.
	Red bone marrow
	5 rems

	3.
	Whole body radiation exposure
	5 rems

	
	
	

	4.
	Skin
	30 rems

	5.
	Bone
	30 rems

	6.
	Thyroid
	30 rems

	
	
	

	7.
	Hands and forearms
	75 rems

	8.
	Feet and ankles
	75 rems

	9.
	Any other single organ (including lens of the eye)
	15 rems


Schedule 4
Classification of radioactive substances into groups for transport purposes
sections 38 and 39

	First Column
	Second Column

	Nuclide
	Group

	Actinium 227
	I

	Actinium 228
	II

	Americium 241
	I

	Americium 243
	I

	Antimony 122
	III

	Antimony 124
	III

	Antimony 125
	III

	Argon 37
	III

	Argon 41
	III

	Arsenic 73
	III

	Arsenic 74
	III

	Arsenic 76
	III

	Arsenic 77
	III

	Astatine 211
	III

	Barium 131
	III

	Barium 140
	III

	Berkelium 249
	II

	Beryllium 7
	III

	Bismuth 206
	III

	Bismuth 207
	III

	Bismuth 210
	III

	Bismuth 212
	III

	Bromine 82
	III

	Cadmium 109
	III

	Cadmium 115m
	III

	Cadmium 115
	III

	Caesium 131
	III

	Caesium 134m
	III

	Caesium 134
	III

	Caesium 135
	III

	Caesium 136
	III

	Caesium 137
	III

	Calcium 45
	III

	Calcium 47
	III

	Californium 249
	I

	Californium 250
	I

	Californium 252
	I

	Carbon 14
	III

	Cerium 141
	III

	Cerium 143
	III

	Cerium 144
	II

	Chlorine 36
	III

	Chlorine 38
	III

	Chromium 51
	III

	Cobalt 57
	III

	Cobalt 58m
	III

	Cobalt 58
	III

	Cobalt 60
	III

	Copper 64
	III

	Curium 242
	II

	Curium 243
	I

	Curium 244
	I

	Curium 245
	I

	Curium 246
	I

	Dysprosium 165
	III

	Dysprosium 166
	III

	Erbium 169
	III

	Erbium 171
	III

	Europium 152 (9.2hr)
	III

	Europium 152 (13yrs)
	III

	Europium 154
	II

	Europium 155
	III

	Fluorine 18
	III

	Gadolinium 153
	III

	Gadolinium 159
	III

	Gallium 72
	III

	Germanium 71
	III

	Gold 196
	III

	Gold 198
	III

	Gold 199
	III

	Hafnium 181
	III

	Holmium 166
	III

	Hydrogen 3
	III

	Indium 113m
	III

	Indium 114m
	III

	Indium 115m
	III

	Indium 115
	III

	Iodine 126
	III

	Iodine 129
	III

	Iodine 131
	III

	Iodine 132
	III

	Iodine 133
	III

	Iodine 134
	III

	Iodine 135
	III

	Iridium 190
	III

	Iridium 192
	III

	Iridium 194
	III

	Iron 55
	III

	Iron 59
	III

	Krypton 85m
	III

	Krypton 85
	III

	Krypton 87
	III

	Lanthanum 140
	III

	Lead 203
	III

	Lead 210
	II

	Lead 212
	III

	Lutecium 177
	III

	Manganese 52
	III

	Manganese 54
	III

	Manganese 56
	III

	Mercury 197m
	III

	Mercury 197
	III

	Mercury 203
	III

	Molybdenum 99
	III

	Neodymium 144
	III

	Neodymium 147
	III

	Neodymium 149
	III

	Neptunium 237
	I

	Neptunium 239
	III

	Nickel 59
	III

	Nickel 63
	III

	Nickel 65
	III

	Niobium 93m
	III

	Niobium 95
	III

	Niobium 97
	III

	Osmium 185
	III

	Osmium 191m
	III

	Osmium 191
	III

	Osmium 193
	III

	Palladium 103
	III

	Palladium 109
	III

	Phosphorus 32
	III

	Platinum 191
	III

	Platinum 193
	III

	Platinum 193m
	III

	Platinum 197m
	III

	Platinum 197
	III

	Plutonium 238
	I

	Plutonium 239
	I

	Plutonium 240
	I

	Plutonium 241
	II

	Plutonium 242
	I

	Polonium 210
	II

	Potassium 42
	III

	Praseodymium 142
	III

	Praseodymium 143
	III

	Promethium 147
	III

	Promethium 149
	III

	Protactinium 230
	II

	Protactinium 231
	I

	Protactinium 233
	III

	Radium 223
	II

	Radium 224
	II

	Radium 226
	I

	Radium 228
	II

	Radon 222
	III

	Radon 221
	II

	Rhenium 183
	III

	Rhenium 186
	III

	Rhenium 187
	III

	Rhenium 188
	III

	Rhenium natural
	III

	Rhodium 103m
	III

	Rhodium 105
	III

	Rubidium 86
	III

	Rubidium 87
	III

	Rubidium natural
	III

	Ruthenium 97
	III

	Ruthenium 103
	III

	Ruthenium 105
	III

	Ruthenium 106
	III

	Samarium 147
	III

	Samarium 151
	III

	Samarium 153
	III

	Scandium 46
	III

	Scandium 48
	III

	Selenium 75
	III

	Silicon 31
	III

	Silver 105
	III

	Silver 110m
	III

	Silver 111
	III

	Sodium 22
	III

	Sodium 24
	III

	Strontium 85m
	III

	Strontium 85
	III

	Strontium 89
	III

	Strontium 90
	II

	Strontium 91
	III

	Strontium 92
	III

	Sulphur 35
	III

	Tantalum 182
	III

	Technetium 96m
	III

	Technetium 96
	III

	Technetium 97m
	III

	Technetium 97
	III

	Technetium 99m
	III

	Technetium 99
	III

	Tellurium 125m
	III

	Tellurium 127m
	III

	Tellurium 127
	III

	Tellurium 129m
	III

	Tellurium 129
	III

	Tellurium 131m
	III

	Tellurium 132
	III

	Terbium 160
	III

	Thallium 200
	III

	Thallium 201
	III

	Thallium 202
	III

	Thallium 204
	III

	Thorium 227
	II

	Thorium 228
	I

	Thorium 230
	I

	Thorium 231
	III

	Thorium 232
	III

	Thorium 234
	III

	Thorium natural
	III

	Thulium 170
	III

	Thulium 171
	III

	Tin 113
	III

	Tin 125
	III

	Tungsten 181
	III

	Tungsten 185
	III

	Tungsten 187
	III

	Uranium 230
	II

	Uranium 232
	II

	Uranium 233
	II

	Uranium 234
	II

	Uranium 235
	III

	Uranium 236
	II

	Uranium 238
	III

	Uranium natural
	III

	Vanadium 48
	III

	Xenon 131m
	III

	Xenon 133
	III

	Xenon 135
	III

	Ytterbium 175
	III

	Yttrium 90
	III

	Yttrium 91m
	III

	Yttrium 91
	III

	Yttrium 92
	III

	Yttrium 93
	III

	Zinc 65
	III

	Zinc 69m
	III

	Zinc 69
	III

	Zirconium 93
	III

	Zirconium 95
	III

	Zirconium 97
	III


Schedule 5
Maximum amounts of radioactive substances exempted from this Act
section 5

	GROUP I  :
	Maximum allowed :
	1 microcurie

	
	Lead (210)
	Radium (226)

	
	Actinium (227)
	Plutonium (230)

	
	Americium (241)
	Polonium (219)

	
	Astatine (211)
	Uranium (233)

	GROUP II :
	Maximum allowed:
	10 microcuries

	
	Scandium (46)
	Cobalt (60)

	
	Strontium (90)
	Ruthenium (106)

	
	Silver (105)
	Tellurium (129)

	
	Iodine (131)
	Caesium (137)

	
	Cerium (144)
	Europium (154)

	
	Tungsten (181)
	Rhenium (183)

	
	Iridium (192)
	

	GROUP III:
	Maximum allowed:
	100 microcuries

	
	Phosphorus (32)
	Chlorine (36)

	
	Calcium (45)
	Scandium (48)

	
	Vanadium (48)
	Iron (59)

	
	Zinc (65)
	Gallium (72)

	
	Arsenic (76)
	Rubidium (86)

	
	Strontium (89)
	Yttrium (91)

	
	Niobium (95)
	Technetium (96)

	
	Rhodium (105)
	Silver (111)

	
	Cadmium (109)
	Tin (113)

	
	Tellurium (127)
	Barium (140)

	
	Lanthanum (140)
	Praesodymium (143)

	
	Samarium (151)
	Holmium (166)

	
	Thulium (170)
	Lutecium (177)

	
	Tantalum (182)
	Platinum (191)

	
	Gold (198)
	Thallium (200)

	
	Thallium (204)
	Lead (203)

	
	Thorium (234)
	

	GROUP IV :
	Maximum allowed :
	1000 microcuries

	
	Hydrogen (3)
	Beryllium (7)

	
	Carbon (14)
	Sodium (24)

	
	Sulphur (35)
	Potassium (42)

	
	Chromium (51)
	Manganese (56)

	
	Iron (55)
	Nickel (59)

	
	Copper (64)
	Germanium (71)

	
	Molybdenum (99)
	Palladium (103)

	
	Promethium (147)
	Iridium (190)

	
	Gold (196)
	Thallium (201)

	
	Thallium (202)
	


NOTES

1.
Figures immediately following a substance in this Schedule refer to the atomic mass numbers of the substance, thus Hydrogen (3) means Hydrogen of Atomic Mass 3.

2.
Any radioactive substance not specified in this Schedule shall be taken to be in Group II.

Schedule 6
Specifications for packages
sections 36 and 37

PART I

DESIGN
1.
The smallest dimension for any outer container for a radioactive substance is 10 centimetres.

2.
A package shall be so designed that it can be easily handled and be properly secured in or on the conveyance during transport.

3.
A package of gross mass of 10 kilograms or more up to 50 kilograms shall be provided with means for manual handling.

4.
A package of gross weight exceeding 50 kilograms shall be so designed as to enable safe handling by mechanical means.

5.
A design does not comply with rules 3 and 4 unless any lifting attachment on the package, when used in the intended manner, does not impose unsafe stresses on the structure of the package.

6.
The outer layer of the package shall be such that it does not collect and retain water.

7.
The external surfaces of the package shall be designed so that they may be easily decontaminated.

8.
The maximum value of radiation originating from a package shall not exceed 200 millirems per hour at any point on the surface of the package or 10 millirems per hour at a distance of 1 metre from the package.

9.
A package shall incorporate a seal on the outside which is not readily breakable and which, while intact, indicates that the package has not been opened.

10.
A package shall be so designed to withstand the effects of any acceleration, vibration and vibration resonance which may arise during normal transport without any deterioration in the effectiveness of the closing devices in the various receptacles or in the integrity of the package as a whole.

11.
The nuts, bolts and other securing devices shall be so designed as to prevent them from becoming loose or being released unintentionally, even after repeated use.

12.
A package shall include in its design a containment system securely closed by a positive fastening device which cannot be opened unintentionally or by pressure which may arise within the package.

13.
A package of which the outer component of a containment system forms a separate unit shall be capable of being securely closed by a fastening device that is independent of any other part of the package.

14.
The materials of a package and its components or structures shall be physically and chemically compatible with each other and with the contents of the package.

15.
The design of any component of the containment system shall take into account any likely radiolytic decomposition of liquids and other vulnerable materials and the generation of gas by chemical reaction and radiolysis.

16.
The containment system shall retain its radioactive contents under the reduction of ambient pressure to 0.25 kilograms per square centimetre.

17.
All valves other than pressure relief valves through which the radioactive contents could otherwise escape shall be protected against unauthorized operation and shall be provided with an enclosure to retain any leakage from the valve.

18.
A radiation shield which encloses a component of a package specified as part of the containment system shall be so designed as to prevent the unintentional release of the component from the shield.  Where the radiation shield and the component within it forms a separate unit, the radiation shield shall be capable of being securely closed by a fastening device which is independent of any other package structure.

19.
A tie-down attachment on a package shall be so designed that, under both normal and accident conditions, the forces in those attachments will not impair the ability of the package to meet the requirements of this Schedule.

PART II

LABELLING AND MARKING
20.
A package and a container shall display at least 2 appropriate labels as prescribed in this Part of this Schedule according to the category into which the package or the container falls.

21.(1)
A package or a container falls into Category I when the radiation level originating from the package or any package within the container at any time during normal transport does not exceed 0.5 millirem per hour at any location on the external surface of the package.

(2)
The appropriate label for packages falling in Category I:
(a)
shall have a white background;

(b)
shall be in the form of Figure 2 in Schedule 2;

(c)
shall display the colours indicated in that Schedule for the lettering and designs in that figure; and

(d)
shall comply with the dimensions indicated in that Schedule for that figure.

22.(1)
A package falls into Category II when the radiation level originating from the package, at any time during normal transport, exceeds 0.5 millirem per hour but does not exceed 50 millirems per hour at any location on the external surface of the package and the maximum radiation level does not exceed one millirem per hour at one metre from the external surface of the package.

(2)
A container falls into Category II when it contains at least one package falling into Category II and the maximum level of radiation does not exceed one millirem per hour at a distance of one metre from the container.

(3)
The appropriate label for packages and containers falling into Category II:
(a)
shall have a yellow background;

(b)
shall be in the form of Figure 3 in Schedule 2;

(c)
shall display the colour indicated in that Schedule for the lettering and designs in that figure; and

(d)
shall comply with the dimensions indicated in that Schedule for that figure.

23.(1)
A package falls into Category III when the radiation level originating from the package at any time during normal transport exceeds 50 millirems per hour but does not exceed 200 millirems per hour at any location on the external surface of the package.

(2)
A container falls into Category III:
(a)
when at any time during normal transport the radiation level exceeds 1 millirem per hour at a distance of 1 metre from the external surface of the freight container; or

(b)
when the radiation level exceeds 200 millirems per hour but does not exceed 1000 millirems per hour at any location on the external surface of the container and the transportation is carried out under the conditions specified in section 45.

(3)
The appropriate label for Category III packages and containers:
(a)
shall have a yellow background;

(b)
shall be in the form of Figure 3 in Schedule 2;

(c)
shall display the colour indicated in that Schedule for the lettering and designs in that figure; and

(d)
shall comply with the dimensions indicated in that Schedule for that figure.

24.
Labels shall be fixed on 2 opposite sides of the outside of the package or container.

25.
Each label shall be completed before despatch by marking on it where indicated in the figures in the Schedule the information relating to the contents and the radioactivity of the contents.

26.
A package of gross mass exceeding 50 kilograms shall have its gross mass plainly and durably marked outside of the package.

27.(1)
A package that conforms to a type A package design shall be plainly and durably marked on the outside of the package "Type A".

(2)
A package that conforms to a type B package design shall be plainly and durably marked "Type B" and shall have the outside of the outermost receptacle which is resistant to fire and water plainly marked (by embossing, stamping or by using some other means of producing a mark resistant to fire and water) with a symbol in the form of Figure 1 in Schedule 2.

28.
A package containing a radioactive substance which has additional hazardous characteristics shall carry additional labels to indicate those characteristics.

ENDNOTES
1
KEY
Key to abbreviations

amd = amended 



od = order

app = appendix



om = omitted

bl = by-law




pt = Part

ch = Chapter




r = regulation/rule

cl = clause




rem = remainder

div = Division




renum = renumbered

exp = expires/expired



rep = repealed

f = forms




s = section

Gaz = Gazette




sch = Schedule

hdg = heading




sdiv = Subdivision

ins = inserted




SL = Subordinate Legislation

lt = long title




sub = substituted

nc = not commenced





2
LIST OF LEGISLATION

	Radiation (Safety Control) Ordinance 1978 (Act No. 87, 1978)

	Assent date
	23 August 1978

	Commenced
	1 February 1980 (Gaz G4, 25 January 1980, p 7)

	Transfer of Powers (Health) Act 1978 (Act No. 122, 1978)

	Assent date
	21 December 1978

	Commenced
	1 January 1979 (s 2)

	Statute Law Revision Act (No. 2) 1979 (Act No. 128, 1979)

	Assent date
	15 October 1979

	Commenced
	15 October 1979

	Statute Law Revision Act (No. 2) 1982 (Act No. 54, 1982)

	Assent date
	8 October 1982

	Commenced
	8 October 1982

	Criminal Law (Regulatory Offences) Act 1983 (Act No. 68, 1983)

	Assent date
	28 November 1983

	Commenced
	1 January 1984 (s 2, s 2 Criminal Code Act 1983 (Act No. 47, 1983), Gaz G46, 18 November 1983, p 11 and Gaz G8, 26 February 1986, p 5)


3
LIST OF AMENDMENTS

lt
amd No. 54, 1982, s 2

ss 1 – 2
amd No. 54, 1982, s 2

s 3
amd No. 122, 1978, s 77; No. 54, 1982, s 2

s 4
amd No. 54, 1982, s 2

s 5
amd No. 122, 1978, s 78; No. 54, 1982, s 2

s 6
amd No. 122, 1978, s 78


sub No. 128, 1979, s 28

ss 7 – 8
amd No. 122, 1978, s 78; No. 54, 1982, s 2

ss 9 – 10
amd No. 54, 1982, s 2

s 11
amd No. 122, 1978, s 78

s 12
amd No. 122, 1978, s 78; No. 54, 1982, s 2

s 14
amd No. 54, 1982, s 2

s 15
amd No. 122, 1978, s 78; No. 54, 1982, s 2

s 16
amd No. 122, 1978, s 78

s 17
amd No. 122, 1978, s 78; No. 54, 1982, s 2

ss 19 – 20
amd No. 122, 1978, s 78

s 22
amd No. 122, 1978, s 78; No. 54, 1982, s 2

ss 24 –27
amd No. 122, 1978, s 78

s 28
amd No. 54, 1982, s 2

ss 30 – 35
amd No. 122, 1978, s 78

s 36
amd No. 54, 1982, s 2

s 37
amd No. 122, 1978, s 78

s 38
amd No. 54, 1982, s 2

s 40
amd No. 122, 1978, s 78

s 45
amd No. 122, 1978, s 78; No. 54, 1982, s 2

ss 47 – 48
amd No. 122, 1978, s 78

s 49
amd No. 122, 1978, s 79

s 50
amd No. 122, 1978, s 78

s 51
amd No. 122, 1978, ss 78 and 80

ss 52 – 53
amd No. 122, 1978, s 78

s 54
amd No. 54, 1982, s 2

s 54A
ins No. 68, 1983, s 28

s 55
amd No. 122, 1978, s 78; No. 54, 1982, s 2

s 56
amd No. 122, 1978, s 81; No. 54, 1982, s 2

PAGE  
Radiation (Safety Control) Act
iii

